## Eligibility Criteria
ELIGIBILITY CRITERIA
| Name of the Hospital | SRMC - Kancheepuram
| Copy of Application Form Verified |  | Maximum Marks | Scored Marks
NO | Criteria
1 | Access | Signages | 1 | 1
|  | CMCHIS boards displaying the available services under the scheme | 1 | 1
2 | Ramp | Availability of Ramp for stretchers / Wheel Chairs | 1 | 1
|  | Ramp broad enough to facilitate 1 stretcher and a wheel chair simultaneously | 1 | 1
3 | Lift | Fail safe incase of power outage (automatically stops and opens in the floor below) & Fire Safety | 1 | 1
|  | Lift wide enough to accommodate stretchers and wheel chairs | 1 | 1
4 | Maintanance | 24 x 7 | 1 | 1
|  | Alternate power source | 1 | 1
|  | Alternate source and back up-for LSS ( vaccum, medical gasses) must check the sorurces and routes | 1 | 1
5 | Fire Safety | Certificate | 1 | 1
|  | Mock drills register, fire safety equipments and supplies in each floor, Operating theatre, Casualty, Lab and ICU | 1 | 1
6 | IMS | OP and IP registry as per the provided criteria (Diagnosis , Investigations,specialists involved, drug charts) must be available electronically | 1 | 1
|  | EHR with unique ID | 2 | 2
7 | BMW | BMW management approval /Renewal at prescribed intervals with Pollution Control Board or Authorized agencies | 1 | 1
|  | Detailed plan of regular disposal and Quantum disposed (bills) for minimum 3 months 0.5 kg estimate per bed occupancy | 1 | 1
8 | Labaratory services | EQAS/ IQAS certified | 1 | 1
|  | ABG | 1 | 1
|  | Outsourced investigations must be with EQAS/IQAS certified centres only ( agrrement copy to be obtained) | 1 | 1
|  | Patho lab (with Pathologist) | 1 | 1
|  | Bio-Chem lab with Autoanalyser, Ultra centrifuge, Electrophoresis, Immunoassays, Elisa | 1 | 1
|  | Micro lab - ( M.Sc - Microbiologist) | 1 | 1
9 | Imaging | PACS database with back up minimum 3 month | 1 | 1
Imaging database includes all CT / X-ray /MRI / USG / ECHO
10 | Pharmacy | Pharmacy license registration certificate | 1 | 1
|  | MIS (computerised) Stock register including expiry data, order sheets, returns for minimum 3 months | 1 | 1
11 | Ambulance | Mandatory oxygen and BLS | 1 | 1
|  | Certified Paramedics round the clock | 1 | 1
|  | RC book, chasis number (to be collect from hospital) | 1 | 1
12 | MRD | Minimum one years files, charts , reports and copies of discharge summary/ Death certificates | 1 | 1
| separate MRD for CMCHIS cases is not acceptable | All files with ICD code, | 1 | 1
13 | HIC | Reports - Ward swabs, | 1 | 1
|  | Reports - ICU swabs/Ventilator Swabs/RO | 1 | 1
|  | Reports - cssd certificate | 1 | 1
14 | OT | Sterility record - certificate | 1 | 1
|  | OT Register | 1 | 1
15 | Ward | Bed Count (EDC may decide)
|  | upto 30 | 1 | 0
|  | 30-100 | 2 | 0
|  | >100 | 3 | 3
16 | Blood Bank | Registered in house blood bank | 1 | 1
|  | with in house Storage facility - functional with 1 year register | 1 | 1
|  | TOTAL | 40 | 40

## Common Score Sheet
NO | Criteria | Common for all specialities | Maximum Marks | Marks Scored
1 | HR | Management Structure with Contacts ( e-mail, phone, Mobile) | 1 | 1
|  | Specialists attendance Log minimum 1 year (all specialists certificates, contact numbers and undertaking of other practicing hospitals to be provided) | 1 | 1
|  | Duty Medical Officers attendance Log minimum 1 year (all certificates, contact numbers and undertaking of other practicing hospitals to be provided) | 1 | 1
|  | Medical Officer Requirement as per IPHS Standards Doctor : Beds ratio 1:6 > 50% 3 marks <50% 1 mark | 3 | 3
Nursing Staff
|  | Qualified Nursing staff attendance Log minimum 1 year (all certificates to be collected) | 1 | 1
|  | Requirement as per IPHS Standard Nurse : Beds ratio 1:2 > 50% 3 marks <50% 1 mark | 3 | 3
Laboratory & Imaging
|  | 24 hrs lab in house with all 4 mentitioned facilities | 2 | 2
|  | radiographer attendance Log minimum 1 year (all certificates, contact numbers and undertaking of other practicing hospitals to be provided) | 1 | 1
|  | lab technician (minimum 2 Round the clock) attendance Log minimum 1 year (all certificates, contact numbers and undertaking of other practicing hospitals to be provided) | 1 | 1
Maintanance & Service staff
|  | Round the clock staff with knowledge on alternative sources for Power and BLS, Emergency Exits, plumbing system etc. - 1 year attendance log | 2 | 2
2 | Radiology | Digital X-Ray (With qualified Radiologist) | 1 | 1
| In House | USG (With qualified Sonologist) | 1 | 1
|  | CT (With qualified Radiologist) | 2 | 2
|  | MRI (With qualified Radiologist) | 2 | 2
|  | Echo (with qualified cardiologist) | 2 | 2
3 | ICU | Minimum 10 beds for common ICU OR 5 beds for speciality ICU | 2 | 0
|  | Manned by an Intensivist / General Medicine / Aneasthetist | 2 | 2
|  | ICU Manned by Medical officer in addition to specialist 1: 6 bed ratio | 1 | 1
|  | Ventilator (High Frequency) (Invoice / PO copy to be collected) only functioning ventilator at the time of inspection to be taken into account | 2 | 2
|  | Support staff trained in using ventilator (certificates to be collected) As per IPHS standards 1:1nurse and bed ratio 1:6 paramedical staff trained in BLS and bed ratio > 50% 3 marks <50% 1 mark | 3 | 3
|  | Emergency medicines tray, Laryngoscope, ET tubes, defibirillator | 2 | 2
|  | Functional Multipara monitor beside every bed | 2 | 2
4 | Accident and Emergency care | Fully equiped with Emergency medicines tray, Laryngoscope, ET tubes, Multiparameters, defibirillator, ECG, Suction, O2, etc. | 2 | 2
|  | Total | 40 | 38

## Speciality score sheet
Speciality wise Score Sheet
Sl No | UROLOGY | Maximum marks | Marks Scored
1 | In house Nephrologist / General medicine (Backup) | 2 | 2
2 | Nephroscope | 2 | 2
3 | Pediatric endo urology equipments | 2
4 | Laser | 2
5 | Flexible ureteroscope | 2 | 2
6 | ESWL machine | 2
7 | Renal Transplant facility- (HOTA) | 4
8 | Haemodialysis | 2
9 | ETO sterilizer | 2
| Total | 20 | 6
Sl.No | Ortho & Spine | Maximum marks | Marks Scored
1 | backed up by Neurosurgeon | 2 | 2
2 | in house Neurologist, General Medicine, Cardiologist | 2 | 2
3 | C- Arm | 2 | 2
4 | Laminar flow | 2
5 | Hepa filter | 2
6 | Arthroscopy equipments | 2 | 2
7 | Operating microscope | 2 | 2
8 | Stryker Frame | 2
9 | DVT prophylaxis protocol | 2
10 | Specility ICU | 2
| Total | 20 | 10
Sl.No | Ophthal | Maximum marks | Marks Scored
1 | General - ophthalmologist |  | YES
2 | Vitero - Retinal surgeon | 2 | 2
3 | Operating Microscope | 2
4 | FFA (in house) | 2
5 | Laser (in house) | 2
6 | B scan (in house) | 1
7 | Occular coherence Tomography (in house) | 3
8 | Boyles | 1 | 1
9 | Autoclave | 1 | 1
10 | Corneal transplant | 3
11 | RNFL & | 3
| Total | 20 | 4
Sl.No | ENT | Maximum marks | Marks Scored
1 | Fully equipped operation theatre | 1 | 1
2 | Operating Microscope | 1 | 1
3 | High definition Endoscopy camera | 1 | 1
4 | Video Flexible Endoscopy services including Laryngoscopy, Bronchoschopy | 1 | 1
5 | Cochlear Implant programme | 1
6 | In house speech evalutaion | 1
7 | Audio verbal Therapy facilities (1 year attendance with contact details of patients) | 1
8 | In house hearing Audiology and vestibular testing facilities | 1 | 1
9 | Sound proof Audiology lab | 1 | 1
10 | Brainstem Evoked Response Audiometry | 1 | 1
11 | Oto acoustic emission ( OAE) | 1 | 1
12 | Auditory steady state response audiometry | 1 | 1
13 | Electo cochleo graphy( ECoG) | 1 | 1
14 | Carnio corphography | 1 | 1
15 | Electro nystagmography | 1 | 1
16 | Vestibular evoked myogenic potenial VEMP | 1 | 1
17 | Pure tone Audiometry | 1 | 1
18 | Impedance Audiometry | 1 | 1
19 | Hearing Aid repair and service facilities | 1
20 | Qualified Audiologist | 1
| Total ( to be factored to 10) | 20 | 15
Sl No | REPLACEMENT | Maximum marks | Marks Scored
1 | Backup Specialists Neurologist, General Medicine, Cardiologist | 2
2 | In Hosue Echo, Doppler | 2
3 | Laminar Air Flow | 2
4 | c-arm in theatre | 2
5 | Hepafilter | 2
6 | Arthroscopic Equipment | 2
7 | DVT Protocol | 2
8 | List of surgeries done so far with number of surgeries with Number of Death and Number of revision surgeries with contact details of patients | 2
9 | Post op follow up records of the patients so far operated (Documents for minimum 1 year) with contact details of patients | 2
10 | Physiotherapist in House | 2
| Total | 20
Sl No | Renal transplantation | Maximum marks | Marks Scored
1 | Neurologist, General Medicine, Cardiologist and surgical backup | 2
2 | Transplant ICU | 2
3 | Step down ICU with Emergency medicines tray, Laryngoscope, ET tubes, defibirillator | 2
4 | Dialysis machines - number (functioning) | 2
5 | Dialysis technicians - number with qualifications | 2
6 | renal transplant OT register (apart form scheme )- Xerox of OT reg to be sent | 1
7 | Human Organ Transplant certificate - Renal | 2
8 | List of surgeries done so far with number of surgeries with Number of Death and Number of redo surgeries with contact details of patients | 2
9 | Post op follow up records of the patients so far operated (Documents for minimum 1 year) with contact details of patients | 2
10 | Clinical Pharmacist | 1
11 | In house Lab facility for evaluation/assessments ( HD HLA Matching) | 2
| Total | 20
Sl No | DIALYSIS | Maximum marks | Marks Scored
1 | Types of membranes used- FDA approved | 2
2 | ICU facility in the hospital with Doctor availability in HD ward. | 2 | 2
3 | Physical separation of Seropositive HD ward | 2 | 2
4 | Separate machines for seropositive cases | 2
5 | Nephrologist availability (consultation minimum once a month for every patient)
6 | Dialyser , Tube reuse [Yes/No] | 2
7 | HD machine serology & culture report [quarterly] | 2
8 | Dialysis membranes/tubes storage in Pigeon holes/Closed Boxes(should be inspected the site by inspecting doctor,send his opinion regarding the hygiene of the storage site) | 2
9 | R.O water culture report from a govt approved laboratory (once in evry 3 months) | 2
10 | Total no of machines | 2 | 2
11 | Dialysis Register (Minimum 1 year) | 2 | 2
| TOTAL | 20 | 8
Sl No | Polytrauma Inspection Format | Maximum marks | Marks Scored
1 | Consultant General surgeon with qualification and certificates |  | Yes
2 | Consultant Orthopedician with qualification and certificates |  | Yes
3 | Consultant Plastic surgeon with qualification and certificates | 2 | 2
4 | Consultant Neuro surgeon with qualification and certificates | 2 | 2
5 | Consultant Vascular surgeon with qualification and certificates | 2 | 2
6 | Consultant Cardiothoracic surgeon with qualification and certificates | 2 | 2
7 | Backup Specialists General Medicine, Neurologist, Cardiologist, etc Minimum 3 Specialists | 2 | 2
8 | Fully equiped theatre C- Arm OT- Laminar flow Hepa filter Arthroscopy equipments Operating microscope | 4
9 | In house EEG, ECHO, USG, Digital X-ray | 2
10 | In house CT & MRI | 2
11 | Physiotherapist in House | 2
| TOTAL | 20 | 10
Sl No | Oncology | Maximum marks | Marks Scored
1 | SUPPORTIVE / BACK UP SPECIALISTS minimum 3 | 2 | 2
2 | Speciality ICU | 2 | 2
Surgical Oncology
3 | C-ARM IN THEATRE | 2 | 2
4 | ENDOSCOPES ( INCLUDING, COLONOSCOPES, LARYNGOSCOPES, BRONCHOSCOPES, ETC. | 2
5 | Stereotaxy Frame Orthogonal / Burr Hole Mounted / Arc-Quadrant / Arc-Phantom | 3 | 3
6 | Hybrid OR Intra-operative CT /MRI | 3
7 | Neuronavigation ,3D reconstruction using Functional MRI or DTI / DOT / | 3
8 | Awake brain Surgery with ECOG | 3
| TOTAL | 16 | 9
RADIATION ONCOLOGY
9 | COBALT EXTERNAL BEAM THERAPY | 2
10 | LINAC PHOTON THERAPY MACHINE | 2 | 2
11 | 3D RT | 2 | 2
12 | IMRT ,IGRT | 2
13 | RAPID ARC | 2
14 | AERB Certification | 2 | 2
15 | RADIATION DOSIMETER BADGE | 2
16 | IN HOUSE FACILITIES FOR MAKING MOULDS | 2
| TOTAL | 16 | 10
MEDICAL ONCOLOGY
17 | PAIN MANAGEMENT CLINIC INCLUDING NARCOTIC USAGE PERMISSION | 2
18 | USG GUIDED NERVE/GANGLION BLOCKS | 2
19 | CT/ C ARM GUIDED NERVE/ GANGLION BLOCKS | 4
20 | TERMINAL CARE FACILITY Fully equiped | 4 | 4
21 | PSYCHOLOGIST COUNSELLING | 2
22 | GROUP THERAPY SESSIONS | 2
| TOTAL | 16 | 8
Sl No | LIVER TRANSPLANT | Maximum marks | Marks Scored
1 | Surgical gastroenterologists with number and qualifications |  | YES
2 | Medical gastroenterologist trained with interventional Gastroenterology with number and qualifications | 2 | 2
3 | ENDOSCOPES ( INCLUDING, COLONOSCOPES, LARYNGOSCOPES, BRONCHOSCOPES, ETC. ERCP | 2
4 | in house Neurologist, General Medicine, Cardiologist and surgical backup | 2 | 2
5 | In house Lab facility for evaluation/assessments ( HD HLA Matching) | 2 | 2
6 | Transplant ICU | 2 | 2
7 | Post transplant Step down ICU with Emergency medicines tray, Laryngoscope, ET tubes, defibirillator | 2 | 2
8 | OT register regarding liver transplant (minimum 1 year) | 2
9 | Organ Transplant certificate - Liver | 2
10 | List of surgeries done so far with number of surgeries with Number of Death and Number of redo surgeries with contact details of patients | 2
11 | Post op follow up records of the patients so far operated (Documents for minimum 1 year) with contact details of patients | 2
| TOTAL | 20 | 10
Sl No | GYNAECOLOGY & OBSTETRIC SURGERY | Maximum marks | Marks Scored
1 | LAPROSCOPY EQUIPMENT & MONITOR WITH VIDEO RECORDING FACILITY [MANDATORY] | 2 | 2
2 | WELL EQUIPPED [ICU WITH CENTRAL OXYGEN/SUCTION, MULTIPARA MONITORS, DEFIBRILLATORS, EMERGENCY MEDICAL TRAY]& MANNED ICU | 2 | 2
3 | ICU - NUMBER OF FUNCTIONING VENTILATORS > 2 | 2
4 | Colposcopy, Cardiotocograph, Cryotherapy | 2
5 | In house General Medicine, Cardiologist and surgical backup | 2 | 2
6 | Transvaginal USG | 2
7 | Pathology Lab with tissue storage facility upto 6 months | 2
8 | In house Pathologist | 2
9 | Neonatology back up | 2
10 | Anaesthetic Back up including Epidural | 2
| TOTAL | 20 | 6
Sl No | BONE MARROW & STEM CELL TRANSPLANT | Maximum marks | Marks Scored
1 | Haematologist
2 | Well equipped and manned ICU | 2
3 | Trained staff in and Intensivist trained in BMT/SCT | 2
4 | Isolation ward | 2
5 | Centrifuge ,Cryo freezer | 2
6 | Total Body radiation facility | 4
7 | Back up specialists | 2
8 | List of trasnplantations done so far with number , with Number of Death and Number of redo transplantations with contact details of patients | 2
9 | Post Transplantation follow up records of the patients so far operated (Documents for minimum 1 year) with contact details of patients | 2
10 | In house Lab facility for evaluation/assessments ( HD HLA Matching) | 2
| TOTAL | 20
Sl No | Vascular Surgery | Maximum marks | Marks Scored
1 | In house Conventional Angiogram with fluoroscopy suite, | 2 | 2
2 | Doppler and portable Doppler | 2
3 | Surgical vascular equipment including for micro vascular surgery (vascular clamps, forceps, neddle holders, operating microscope etc..) with fully equiped ICU | 2 | 2
4 | RF abflation machine with relevant disposables/catheters/leads | 2
5 | Post op follow up of the patients so far operated under the scheme. | 2
| TOTAL | 10 | 4*2=8
Multiply the final score by 2
Sl No | Cardio and CTS | Maximum marks | Marks Scored
1 | Cath Lab | 2 | 2
2 | Cardiothoracic OT with Cardiopulmonary by pass equipment | 2 | 2
3 | Perfusionists - number with quailfication certificates | 2 | 2
4 | CCU(Coronary Care Unit) with number of functioning ventilators, central oxygen/suction, multipara monitors, defibrillators, emergency medical tray | 2 | 2
5 | Post OP CTS Ward with number of ventilators | 2 | 2
6 | Facilities for Electrophysiological Studies | 2
| RF abalation and 3D mapping | 2
7 | Facilities for pacemaker implantation | 2
8 | TMT | 1 | 1
9 | Transesophageal ECHO | 2 | 2
10 | Epicardial ECHO | 1 | 1
| TOTAL | 20 | 14
Sl No | Heart and Lung Transplant | Maximum marks | Marks Scored
1 | Cath Lab
2 | Cardiothoracic OT with Cardiopulmonary by pass equipment | 2 | 2
3 | Perfusionists - number with quailfication certificates | 2 | 2
4 | Transplant ICU with number of functioning ventilators, central oxygen/suction, multipara monitors, defibrillators, emergency medical tray | 2 | 2
5 | Step Down ICU with Emergency medicines tray, Laryngoscope, ET tubes, defibirillator | 2 | 2
6 | Facilities for pacemaker implantation | 2
7 | IABP | 2
8 | ECMO | 2
9 | List of trasnplantations done so far with number , with Number of Death and Number of redo transplantations with contact details of patients | 2
10 | Post Transplantation follow up records of the patients so far operated (Documents for minimum 1 year) with contact details of patients | 2
11 | HUMAN ORGAN TRANSPLANT CERTIFICATE | 2
| TOTAL | 20 | 8
Sl No | COCHLEAR IMPLANT & ABI | Maximum marks | Marks Scored
1 | Neuro Surgeon trained in ABI - number with qualifications |  | YES
2 | Minimum 10 surgeries performed as a private player
3 | ABR/ASSR assessment facilities | 2 | 2
4 | OAE assessment facilities
5 | In house CT/MRI - optional
6 | Trained ENT surgeon - number with qualifications | 2
7 | Assistant surgeons- number with qualifications
8 | Trained staffs- number with qualifications
9 | Audiologist- number with qualifications
10 | Speech therapist- number with qualifications
11 | Complete audiology lab with sound proof room | 2 | 2
12 | Pure tone audiometry
13 | Impedance audiometry
14 | BERA
15 | Facilities for Hearing Aid services with technician
16 | AVT Infrastructure AVT specialist-with their qualification and copies of the certificate | 2
17 | Post op follow up records of the patients so far operated with Contact details of the patients | 2
| TOTAL | 10 | 4*2=8
Multiply the final score by 2
Sl No | NICU, PICU, PEDIATRIC AND PEDIATRIC SURGERY | Maximum marks | Marks Scored
1 | Neonatologists [DM Neonatology]number with qualifications | 2 | 2
2 | fully equipped NICU - NICU with central oxygen/suction, multipara monitors, defibrillators, Laryngoscopes, ET tubes and emergency medical tray | 2 | 2
3 | Warmers/incubators > 6 | 2 | 2
4 | Transport incubators | 2 | 2
5 | Ventilators - number of functioning ventilators >3 | 2
| number of functioning ventilators >6 | 3 | 3
6 | PICU with central oxygen/suction, multipara monitors, defibrillators, Laryngoscopes, ET tubes and emergency medical tray | 2 | 2
7 | PICU with number of functioning ventilators 3-5 | 2
| PICU with number of functioning ventilators >6 | 3
8 | pediatric cardiologist practicing at hospital -number with Certificates | 2
9 | pediatric Surgeons practicing - number with certificates | 2 | 2
| TOTAL | 20 | 15
Sl No | NEUROSURGERY | Maximum marks | Marks Scored
1 | Qualified Neurosurgeons : number with certificates |  | YES
2 | Fully equipped theatre with C Arm, operating microscope, endoscopy for spine and display monitors | 2 | 2
3 | stereotactic frame Orthogonal / Burr Hole Mounted / Arc-Quadrant / Arc-Phantom | 2 | 2
4 | Laminar airflow | 2
5 | Hepa-filter
6 | EEG, EMG | 2 | 2
7 | Record of Skull base surgeries minimum 10 | 2
8 | Qualified neurophysiologists number with certificates | 2
9 | In house 24 hr CT /MRI | 2
10 | Hybrid OR Intra-operative CT /MRI | 3
11 | Neuronavigation - 3D reconstruction using Functional MRI or DTI / DOT / Awake brain Surgery with ECOG | 3
| Total | 20 | 6
Sl No | General Surgery, Plastic surgery, Surgical Gastroenterology | Maximum marks | Marks Scored
1 | Surgical Gastro Enterologist |  | YES
2 | General Surgeons |  | YES
3 | Facilities for Information gastro Enterology Procedures including ERCP, Esophageal, Rectal / Anal Manometry Manometry | 2
4 | ENDOSCOPES ( INCLUDING, COLONOSCOPES, LARYNGOSCOPES, BRONCHOSCOPES, ETC. | 2
5 | tissue expanders, automated skingraft Harvest and pin punching machines and plasty equipmets | 2 | 2
6 | Operating Microvascular Microscopes | 2 | 2
7 | Burns ward with Physiotherapy / Rehabilitation | 2 | 2
| TOTAL | 10 | 6*2=12
Multiply the final score by 2
Sl No | Medical specialities including MGE | Maximum marks | Marks Scored
1 | Surgical Gastro Enterologist |  | YES
2 | Medical Gastro Enterologist |  | YES
3 | ERCP, OGD/Colonoscopy/Sigmoidoscope/Isophageal Mamomentry/Rectal / Anal Mamomentry | 2
4 | EEG, EMG, Doppler | 2
5 | electro cautery/ cryotherapy/ uv therapy | 2
6 | Procedure room for Biopsies, scopies etc, with ETO | 2
7 | Back up :Surgical ICU - ventilator support with number of functioning ventilators/ Emergency medicines tray, Laryngoscope, ET tubes, defibirillator/ICU doctors trained in ventilator care/Intensivist | 2
| TOTAL | 10
Sl No | Pulmonology Inspection Format | Maximum marks | Marks Scored
1 | PUMONARY FUNCTION TESTING
2 | IN- HOUSE :Complete PFTs + Pre & Post Bronchodilator Testing (spirometry, lung volumes and DLCO & albuterol 2.5 mg nebulizer) Spirometry (includes flow-volume loop and MVV) , Complete PFTs (spirometry, lung volumes and DLCO) , Lung Volumes (TLC, FRC, RV, VC, etc.) , Maximum Inspiratory and Expiratory Pressures (MIP / MEP ), Carbon Monoxide Diffusing Capacity | 2
3 | FUNCTIONAL CAPACITY
4 | 6 Minute Walk Test (following ATS standards) | 2
5 | Multiple Pulse Oximetry Test
6 | Multiple Pulse Oximetry Test with Oxygen Titration
7 | CardioPulmonary Exercise Testing- Sub-maximal (Step)
8 | CardioPulmonary Exercise Testing- Maximal (Stationary Bicylcle)
ADDITIONAL TESTING
9 | Methacholine Challenge Testing (Methacholine testing, includes FENO) | 2
10 | Arterial Blood Gas | 2
11 | Exercise-induced Bronchospasm | 2
12 | Over Night Oximetry Study
13 | Resting/walking pulse oximetry (non invasive measurement for O2 qualification)
14 | 24 hour Oximetry Study
15 | Fractional Exhaled Nitric Oxide (FENO)
16 | Hemoglobin (Hgb), transcutaneous
| TOTAL | 10
Multiply the final score by 2